# Installing manual

« Les Espaces de Balma »

18 avenue Charles de Gaulle  - Bât. 34 - 31130 BALMA

Tel :33 (0) 562 571 571

www.medicapteurs.fr

Foreword

You are a new MEDICAPTEURS product’s owner of and we thank you.

The MEDICAPTEURS platform is a high performance system. Such an exceptional quality of results demands a great care of the platform. It’s essential to read this manual before using the platform.

The MEDICAPTEURS software was created to be « User-Friendly », but this shouldn’t keep you from reading the user manual.

Added to this manual, you can find the platform and software User Manual on the Medicapteurs CDRom  or USB key.

ENJOY READING THE USER MANUAL AND WELCOME TO THE MEDICAPTEURS PLATFORM USER CLUB.

Quality requirements

This product was tested and approved for the Annex VI, section 3 of the Directive 93/42/EEC for Medical Device.

DEKRA Certification GmbH (Identification Number 0124) approved the Quality System used for Products Final Test.

MEDICAPTEURS system is in accordance with these following norms :

- EN 60601-1-1 August 2001

- EN 60601-1-2 August 2005

- EN 60601-1-4 October 1999

- EN 60601-1-6 September 2007

- EN 62366 March 2008

- EN 62304 October 2006

- EN ISO 14971 November 2009

Medicapteurs is certified for its quality system.

Summary

Foreword	3

Quality requirements	4

SUMMARY	5

Package content	7

available and spare parts	8

Win-pod platform	8

Win-track platform	10

Platform presentation	11

Technical information	11

Medical purpose	13

Indications	13

Contraindications	13

Side effects	13

Hygiene/cleaning	14

Logos	14

Instructions	15

software and driver installation	16

required configuration 	16

Win-pod WIFI : WiFi adapter installation 	16

installation	17

Win-pod WiFi use	20

Turn the platform on 	20

Operation indicator	21

Connecting the wired Win-pod	22

Connecting the platform 	22

Operation indicator	22

Connecting the Win-track	23

Connecting the platform	23

Peripheral connections	24

Operation indicator	24

Scanner	25

Technical information	25

Scanner lock	26

Connecting the scanner	27

Bluetooth remote installation and configuration	28

Computer configuration	28

Mobile configuration	30

software registration procedure	31

Licence checking 	31

Licence registration procedure	31

medicapteurs database import	33

using the platform	34

Static acquisition	34

Dynamic acquisition	34

Postural acquisition	34

Procedure to shut down the platform	35

Package content

Check that your package contains:

- Platform

- Wire

- Installing and using manuals

- Installing CDRom or USB key

- For Win-pod WiFi platforms : two batteries with their USB cable and a charger, a WiFi key.

- For Win-track platforms : a key, an USB connection cover, two screws and a jack reprogramming in a plastic bag

LIMITATION ON WARRANTY

The platform is delivered in a customized packaging to protect the sensitive parts (electronics and sensors).

In case of after sale service, the warranty will not be applied if the product is not shipped back in its original packaging.

We recommend you to keep the original packaging even after the warranty period.

If needed, Medicapteurs will invoice a new packaging for you product shipment.

Available and spare parts

Win-pod platform

| Suitcase (Ref :ValiseWinPodFusyo) |  |
| --- | --- |
| Light suitcase (Ref : Valise_T_Plate) |  |
| Positioner (Ref :PositionneurWinpod) |  |
| For wired platforms : 4.5 meters wire (Ref :CableWPod4_5) |  |

Available and spare parts

| For WiFi platforms : Silver batterie (Ref : Powerblank1silver) |  |
| --- | --- |
| For WiFi platforms : Black batterie (Ref : Powerblank1noir) |  |
| For WiFi platforms : USB charger (Ref : Transformateur10) |  |
| For WiFi platforms : Wifi key (Ref : Cle8) |  |

Available and spare parts

Win-track platform

| 5 meters wire (Ref : CableWT_5) |  |
| --- | --- |
| Light suitcase (Ref : ValiseWintrackLogo) |  |

# Platform presentation

Technical information

| General |
| --- |
| Sensor type | Resistive |
| Min/max pressure range per sensor | from 0,4 to 100 N |
| Acceptable temperature (operation and storage) | from 0 to +40°C |
| Humidity | from 30 to 80 % HR |
| Power surge category | I |
| Acquisition frequency | Up to 200 images/sec |
| Accuracy | 5% |

| Win-pod |
| --- |
| Active surface | 400 x 400 mm |
| Sensor size | 8 x 8 mm |
| Sensor number | 2304 soit 48 x 48 |
| Wired Win-pod |
| Size (Length/Width/Heigth) | 600 x 530 x 45 mm |
| Weight | 6,6 kg |
| Power | via USB (5V, 400mA) |
| PC/platform connection type | USB |
| Win-pod WiFi |
| Size (Length/Width/Heigth) | 610 x 530 x 25 mm |
| Weight | 6.8 kg |
| Power | High capacity rechargeable battery |
| PC/platform connection type | Wifi or USB |

# Platform presentation

Technical information

| Win-track |
| --- |
| Size (Length/Width/Heigth) | 1610 x 652 x 30 mm |
| Weight | 11,8 kg |
| PC/platform connection type | USB |
| Power | via USB (5V, 400mA) |
| Active surface | 1500 x 500 mm |
| Sensor size | 7,8 x 7.8 mm |
| Sensor number | 12288 (192 x 64) |

# Platform presentation

Medical purpose

This is a podometry platform which studies the foot pressure mapping and the orthostatic balance in static, dynamic or postural acquisition. It allows you to objectively assess biomechanical parameters of the foot for the pathology diagnostic process.

This platform can be used by many health specialties: Podiatry, Physical therapy, Sport Medicine, Rehabilitation, Orthopaedics, Rheumatology, Chiropractic…

Indications

Read the following instructions to install and use the system correctly.

Contraindications

Monitor the patients with severe balance problems while making acquisitions. During some protocols, you need to stay close to keep the patient from falling.

In order to be in accordance with the EN 60601-1 norm, current version, the platform doesn’t have to be disassembled.

Side effects

There are no known side effects.

The product development takes into consideration skin allergy risks and electric shock hazard.

# Platform presentation

Hygiene/Cleaning

- After each patient, clean the platform’s upper surface with a product like «cleaning wipes» or any type of antiseptic wipes for medical devices. Do not use disinfectant spray containing alcohol.

- For a best hygiene quality, the patient must wear socks; the measure alterations are considered negligible.

Logos

: Medical Device in accordance with the ECC 93/42 Directive.

: DEKRA Certification Organism

0124		 : Dekra Official Recording Number

: USB connection Device

: « B » type Device

: WEEE device (Medical Electric and Electronic Equipment).

Do not trash. Medicapteurs is in charge of its Waste recycling. The black rectangle means that the plaform was built after 8th August 2005.

: Incentive for user to read instructions.

# Platform presentation

READ THIS BEFORE USING THE PLATFORM

The platform is a very sensitive and accurate tool. The sensor quality response depends on the respect of the following instructions.

Instructions

- The sensors’ life time depends on the respect of those indications. Be careful with the platform when you move it.

- Please handle the platform with caution as you move it, and avoid shocks on the sensors. If you have to ship the platform, don’t forget to protect the active area.

- Always use the same USB port on the PC (do not change the USB Plug-in).

- The platform must be plugged exclusively into a normalized computer EN 60601-1, current version (particularly as it pertains to power connection). No peripheral device should be connected to the computer except if it is self powered (key board, mouse, USB key) or in compliance with regulations EN 60601-1, current version.

- Take care not to bend the connectors of the platform and to unplug the USB from the PC cable prior to any manipulation of the platform.

- Save your database regularly.

- In order to be in accordance with the EN 60601-1 norm, the platform doesn’t have to be disassembled.

- Do not expose the platform directly to the sun (risks of burns).

- This product doesn't need any specific care about the EMC (electromagnetic compatibility). It is an EN 60601-1, current version, certified product so it has been tested by a qualified official organism (Emitech).

- To get a correct acquisition, the platform must be laid on a hard, flat, and dry ground.

- If you stand on the platform with shoes on, check if there might be something dangerous for the platform (like stone, stud...). Do not wear high heels while on it.

- The patient must not step barefoot on the platform. Please provide socks if necessary, to avoid potential reactions due to foot/surface contact.

- The patient must not step on the hood of the platform, and must only put their feet on the sensitive part.

Software and driver installation

WARNING!!! IF you already have a medicapteurs Software ON YOUR COMPUTER, don’t forget to save the database before installing the NEW software.

Required configuration

You may consult us before changing your computer.

To install and use the software, you need a computer including these characteristics:

- Type : PC Dual Core minimum

- Ram : 4 Go minimum

- System : Windows 7, 8 or 10

- At least 2 USB ports dedicated

Win-pod WiFi : WiFi adapter installation

Do not plug the adapter in before installing the drivers.

In the Medicapteurs installation USB key, you will find the setup (as shown below) in the Drivers\WIFI\Setup.exe, that you’ll simply have to run.

Once the installation has ended, you can plug the adapter in. It’s important to always use the same USB port for the adapter.

Software and driver installation

Installation

- Do not connect your platform.

- Power up your computer, or close all applications if it’s already on.

- Insert the software CDrom or the USB key in the reader.

- Double-click on the « Computer » icon then on your CDrom or USB key reader.

- Double-click on this icon        :

- Choose the installation language and click on « OK ».

- Windows displays this screen :

- Click twice on “Next”.

- Choose the platform: select your platform.

- If you have a Win-track platform, go to 11.

If you selected Winpod, choose the platform type: « Win-Pod 2 (USB powered) » or Win-pod WIFI

Software and driver installation

- Choose your measuring system :

- Choose your language :

- If you already have a Medicapteurs’ product, to keep your existing database, choose :

- During this installation, depending of your operating system, the following window could appear, check « Always trust software from Medicapteurs France » and then click on « Install » :

- During installation, a warning from your anti-virus could appear. Click on "Ok" to avoid stopping the installation.

Software and driver installation

- If you have a Winpod WiFi platform : select the Wi-Fi adapter that was given to you with the platform (TP-LINK wireless usb adapter). And Click on Enregistrer.

- Click on « Finish » then « OK » to restart your computer.

# WIN-POD WiFi use

Turn the platform on

- To turn it on, you need to plug the battery in, push-button facing up.

- The green LED will turn on, signaling that the platform is powered.

- If the battery is already plugged in and the LED is turned off, push the button.

# WIN-POD WiFi use

Operation indicator

Initialization

When turning the platform on, a ten-seconds long initialization phase will be represented by the successive flashing of the LEDs within the logo.

Functionning

When the platform is on and in use, the LEDs will flash, as if following each other. The central LED is on.

Stand-By

When the platform is on stand-by, the LEDs are synchronized and pulse together. The platform will enter stand-by mode after 3 minutes of inactivity in the software, and when you close the software.

To exit the stand-by mode, push on the “Power” button on the platform, as shown here :  be sure the green LED on the battery lights up. If not, push the battery’s button.

# Connecting the wired WIN-POD

Connecting the platform

To get an accurate acquisition, the platform must be imperatively laid on a hard and flat ground. The platform’s anti-slid patches will stabilize the platform on a slippery ground during the dynamic acquisition.

- Unplug all USB devices (except mouse and keyboard).

- Take the wire, then plug it in the platform’s connector.

- Screw the blue ring to lock it with the platform.

- Plug the USB connector with a PC USB port (the LED will lit up green).

- VERIFICATION: To check that the driver is correctly installed, go to « Start / Control Panel / System / Hardware / Device Manager / USB controllers ».

The line « Podometer MC board » will appear if the installation was faultless.

Operation indicator

Thanks to the 3 LEDs, you can check the platform:

# Connecting the WIN-TRACK

Connecting the platform

To get an accurate acquisition, the platform must be imperatively laid on a hard and flat ground. The platform’s anti-slid patches will stabilize the platform on a slippery ground during the dynamic acquisition.

- Unplug all USB devices (except mouse and keyboard).

- Insert the USB key into its connector, located on the back of the Win-Track

- Put the cover on the USB connection plug on the Win-Track and use the 2 screws to lock it.

- Connect the other end of the USB cable to the computer. You will later install the Win-Track driver.

- VERIFICATION: To check that the driver is correctly installed go to « Start / Control Panel / System / Hardware / Device Manager / USB controllers ».

The line « PDM MC board » will appear if the installation was faultless.

# Connecting the WIN-TRACK

Peripheral connections

It is possible to connect external peripheral devices to the platform (In and Out connectors located on the back of the platform). These peripheral connectors are designed to synchronize measurements between external devices and the platform (the platform can synchronize the peripheral device or the peripheral device can synchronize the platform).

For example: when Win-Track starts a measurement, it can launch another measurement at the same time on a connected EMG device to correlate measurements with the same timing.

In connector: the external peripheral device synchronizes the platform.

Out connector: the platform synchronizes the external peripheral device.

Devices that can be synchronized: EMG, Video…

All these devices have to be approved by Medicapteurs and must comply with EN 60601-1 regulations (particularly as it relates to power).

Please contact Medicapteurs for specific complementary information prior to using these peripheral connectors.

Operation indicator

Checking function status with the 3 lights on the back of the platform:

# Scanner

Technical information

| General |
| --- |
| Image Sensor | CCD |
| Optical Resolution | 1600 dpi |
| Scan Modes | Color: 48-bit input, 24-bit output Grayscale: 16-bit input, 8-bit output B/W: 1-bit output |
| Color Depth | Input 48-bit; Output 24-bit |
| Scanning Speed | 8.5 sec (Color/ Grayscale/ B&W mode, 300dpi, A3) |
| Scanning Area (W x L) | Max. 304.8 mm x 431.8 mm, A3 |
| Daily Duty Scan ( Sheets ) | 2,500 |
| Acceptable temperature (operation and storage) | from 0 to +40°C |
| Power Supply | 24 Vdc/ 1.67 A |
| Power Consumption | < 36 W (operation), < 4 W (idle) |
| Interface | USB 2.0 |
| Net Weight | 7.5 Kg |
| Dimensions (W x D x H) | 623 x 400 x 133 mm |
| Maximum weight allowed | 120 Kg |

# Scanner

Scanner Lock

Your scanner is designed with two different lock statuses to protect the optical parts. To ensure the scanner working properly, you have to choose an appropriate lock status according to the situation.

Always Unlock:

A scanning-ready status. Furthermore, if you place your scanner on a long–term position, you can also keep your scanner lock in this status. To unlock the scanner, slide the locking switch on the scanner’s underside toward the back panel of the scanner until it snaps into place.

Always Lock:

It holds the scanning head in place during transportation and traveling. The scanner will be kept in disabled status. To lock the scanner, slide the locking switch on the scanner’s underside toward the power LED of the scanner until it snaps into place.

Warning: When you move the scanner, do lock the scanner to ensure accurate and problem-free use of your scanner.

# Scanner

Connecting the scanner

WARNING: Please make sure the scanner is turned off before you plug or unplug the scanner power adapter.

- Unlock your scanner.

- Plug the AC adapter into the scanner’s power receptor and attach the power cable to the AC adapter.

- Plug the other end of the power cable into a standard AC power outlet.

- Plug the square end of the included USB cable to the USB port at the rear of the scanner.

- Plug the rectangular end of the USB cable to an available USB port on your computer.

- If you plan to connect your scanner to a USB hub, make sure that the hub is connected to the USB port in your computer. Next connect the scanner to the USB hub.

# Bluetooth remote installation and configuration

Computer configuration

- Search for « change bluetooth » in « Start / Control panel », then click on « Change Bluetooth settings» :

- Check « Allow Bluetooth devices to find this PC» (or « this computer » depending on your Windows version), then click on « Apply » :

- Go in the « COM Ports » tab and click on « Add » :

# Bluetooth remote installation and configuration

- Choose « Incoming » then click on « OK » :

- Once the device is added, take note of its paired COM port so that you can configure the software:

- Launch the software and go in Configuration \ Parameters \ Plate. Type in the COM port associated with the Bluetooth device:

- In the Acquisition tab, click on  to enable the Bluetooth link from the software. The square indicator next to the button will become green if the connection is correctly made.

# Bluetooth remote installation and configuration

Mobile configuration

This software can be used on a smartphone or tablet:

- With the Android operating system (version 4.1 minimum)

- With Bluetooth

- From your mobile device, go in the Bluetooth settings, look for your computer, then pair them.

- Download and install the Podo Remote app by Medicapteurs from the Google Store.

- Launch Podo Remote with the  icon, then tap on .

- Select your computer then tap on OK. This configuration is saved, and your computer’s name will appear under the  icon.

- Tap on  to open the Bluetooth link from your mobile device. The button’s indicator becomes green if the connection is correctly made.

- Then select, accept or cancel acquisitions from your mobile device.

Software registration procedure

During the software’s first installation, the licence will be automatically registered (in green). You can skip this step.

If, later on, you want to install a protected update (in red), here’s how to register it.

To launch the software, click on the  icon.

You only have 15 days atfer the Installation

to register the licence

After that, you cannot use your

software anymore

Licence checking

On the starting screen, you can see the software’s serial number (25 numbers).

Red background: your software is not registered.

Green background: your software is correctly registered

Licence registration procedure

- Note the software identification numbers.

To find it, click on “Help” then “About” in the software.

- On the screen displayed, note the 25 licence numbers :

Software registration procedure

- Registration :

Please email your Seller so that they can deliver your new serial code.

- Enter your code :

  - Click on « Help » desk then « About »

- In the newly opened window, click on

- Type the new code in this window

- Click on

This screen appears:

Your software is registered and ready to use!

Medicapteurs database import

You can import MEDICAPTEURS database in the software:

To import your old MEDICAPTEURS database

You can find the database in the « data » folder, in the software’s installation folder.

In this sample, we will copy the old PC database on a USB key to import it in the software

- Launch « DbManager » in the software’s installation folder :

- Click on  « Database Import »

- DB Manager displays tables to import

- Click “Next”

- DB Manager displays the data you can import; click on “Begin” to proceed.

Using the platform

Static acquisition

There is no gait reference. The patient only has to stand with their heels on the same horizontal line (it will be displayed on screen).

The patient has to stand on the platform by the rear side.

To get a correct acquisition, arms must be held close to the body and eyes looking straight ahead. Make sure there is nothing in the patient’s pockets.

On screen, you must see two red lines appearing at the rear of the heels. You need to move the patient’s feet to make a single green line.

To record the measure, please read the software User Manual.

Dynamic acquisition

The patient must walk from one of the lateral sides of the platform to the other, by stepping on the active surface. Walk with an even, and natural stride. Walk two steps before putting your lead foot on the platform, then two other steps afterwards.

Postural acquisition

Please follow the same procedure than for the « static acquisition ».

A 30° positioner is available for Win-pod platforms in spare parts to help you for the gait reference.

Posturology exams may be difficult for some patients with balance issues: stay close to them.

Using the platform

Procedure to shut down the platform

To shut down the platform:

- If an acquisition is in progress, just click on Cancel (the large red cross).

- Exit the software (click on the X located on the upper right corner of the window, then click OK).

The platform is on stanby (still connected to the PC).

- To turn off the platform’s “standby” mode:

  - Shut down the computer.

  - Unplug the USB cable linked to the computer.

COPYRIGHT

Copies or reproductions of this document is strictly limited to private use and not authorize for public or collective use, and excerpts or short quotes from this document need to clearly include the author and the source.

All rights reserved. Duplication of any part of the document or the entire document without express written consent of Medicapteurs is illegal and strictly prohibited.

# ©  MEDICAPTEURS  -  Toulouse  -  FRANCE  -  May 2017

# Notes

Notes

« Les Espaces de Balma »

18 avenue Charles de Gaulle  - Bât. 34 - 31130 BALMA

Tel :33 (0) 562 571 571

www.medicapteurs.fr

/ 0124

Copyright 2017

# Text boxes

Red light

Led flashes = Data transfer to the PC

Green light

Led on = Platform powered by the PC

Red light

Led on = Platform recognized by the PC (driver installed)

Light Tf

Light blinking = Transfer of data to PC

Light Ld

Light on = Platform properly connected to PC (driver installed)

Light PW

Light on= Power to the platform